Uncertainty and proficiency testing programs of measurement

of the analysis carbon, silicon, manganese and chrome

in low alloy steel by optical emission spectroscopy OES,

calibrated with a certified reference material

# Dr Elmaghrabi A. M.(1), Chem. Morsey A. S.(2)

(1) ARCOSTEEL Lab Mgr, Sadat City, P.O. 11, Egypt

Email : gmaghraby@yahoo.com, Mob. +20101983932

(2) Lab Sec. Mgr, ARCOSTEEL

# Abstract

The experience of working under an ISO/IEC 17025-2005 compliant quality system in ARCOSTEEL laboratory for special steel analysis, especially uncertainty calculation as a correlation with proficiency test of medium alloy steel by using optical emission spectroscopy OES calibrating with a certified reference material. This study is an example of uncertainty calculation compared with proficiency test study (z factor value) and robust mean standard deviation.

# Introduction

Measurement uncertainty is the most important criterion in both method validation and internal quality control, the increasing implementation of accreditation and quality control pinpoints the need of suitable approached to estimate the measurement uncertainty. This is reflected in the accreditation standard En-ISO/IEC 17025 whose clause 5.4.6.2 states that testing (1) laboratories shall have and shall apply procedure for estimating uncertainty of measurement, proficiency testing is an external quality assurance method in which laboratories are sent simulated specimens for testing by routine laboratory methods, several reports suggests that providing feedback on proficiency testing results improve the quality of testing among different laboratories (2-3). Uncertainty components are a combination of many factors that may include, storage conditions, equipment, an analytical method used, reference, standards, calibration, operator effects and data handling, …etc. the procedure used for estimating the overall uncertainty depends on the data available about the method performance.  According to two components were identified as potential source to the combined standard uncertainty, one source was from the reference laboratory (standard deviation of the results obtained during homogeneity and stability testing) and the second was the reproducibility of the results reported by the participants after removal of outliers, was evaluated in terms of standard deviation of mean of the reported means(4). These components were combined together and combined standard uncertainty was calculated by using the relation.

δc = δ2(hom)+δ2(stab)+δ2(rep)

Where :-

δc 	   is the combined /target uncertainty

δ hom  is the standard deviation of all data collected during homogeneity  study

stab     is the standard deviation of the mean of data collected during homogeneity study

## Analysis principle:-

Measuring system is therefore based on the physical phenomenon that is summarized as follows:

When certain energy is applied to an atom some of its electrons change their orbit, when these electron return to their initial orbit, a precise energy is restored in the form of light at a determined wavelength. A sample containing several different elements will therefore produce light composed of wavelength specific to each of the elements, by separating these wavelength by dispersion system, we can determine which element are present, the intensity of each of these wavelength being a function of the concentration of the considered element by measuring this luminous intensity (with a photo multiplier) and by processing this information with a computer, we can thus determine the concentration of the considered element

## Sample preparation:-

### It is essential that air does not penetrate inside the sparking chamber as air contains  80% N2  and 20% O2. This implies that the sample must be perfectly flat in order to have good air tightness around the analysis table’s hole. This surface condition is obtained by rectifying or milling the sample however, an automatic machine with abrasive belts may also be suitable, even if the surfacing is not ideal, the result is hardly affected The manual preparation of sample often results in a surface more or less convex. This means that air can penetrate into the sparking chamber and seriously deteriorate the analytical results also these condition guarantee the homogeneity of samples.

#### Calculation of uncertainty

#### 1- Repeatability:-

By using optical emission spectroscopy OES 3460 (ARL) has been analyzed the selected sample 10 times on the stabilization of environmental human factors and instrument drift correction

The repeatability calculated according to this formula

q =  1nj=1nqj

S(qj) =  1n−1j=1n(qj−q)2

S(q) = S(qj)n

Where:-

q 		= Mean

S(qj) 	=  standard deviation

S(q) 	= repeatability

Table (1) shows the certified reference sample (CRM) used in uncertainty calculation and the value of every element corresponding to nominal uncertainty, used in these calculation a normal probability distribution with divisor 2 (dividing uncertainty by 2), also repeatability obtained in table according to previous calculation

Table (1) : Nominal uncertainty of CRM with calculated repeatability

| Element | Material Used | Certified Value | Nominal Uncertainty | Calculated Repeatability |
| --- | --- | --- | --- | --- |
| C Si Mn S P Cr | CRM 15259 CRM 187A CRM 1761 CRM 187A CRM BS14B CRM 15259 | 0.67 0.567 0.678 0.018 0.028 0.47 | 0.03 0.023 0.005 0.001 0.002 0.03 | 0.001403 0.001748 0.000616 0.0003 0.000177 0.000904 |

2- CRM standard error

Calculation error is a calculation used in multivariable regression (MVR) during built a calibration curve in OES instrument for used in analysis, this calculation give a value to the difference of the calculated concentration and real concentration, often the procedure of fitting a set of measurement to polynomial equation (up to degree 3) as part of the MVR(6), the calculated curves are then used to determine the concentration of unknown sample, as shown in table (2)

Table (2) : Standard error of curve for each element

| Element | Range |  | SE |
| --- | --- | --- | --- |
|  | Minimum | Maximum |  |
| C Si Mn S P Cr | 0.003 0.004 0.0064 0.015 0.00154 0.0015 | 1.311 1.7 2 0.095 0.111 5.4 | 0.00928 0.00924 0.01368 0.00363 0.00156 0.01853 |

3- Bias from standard samples

We already analyzed the sample over along period (long variation) can calculated the bias as shown in table (3) for every sample, by using the following statistical equations.

MD = 1n(certified value –Measured value)

SD = 1n−1[(certifiedvalue−Measuredvalue)−MD]

Bias from standard samples includes, the uncertainty, environmental variation (temperature, humidity and electromagnetic waves effect, electrostatic charge)

Table (3) : Bias of standard deviation from standard sample

| Element | Material Used | Value |
| --- | --- | --- |
| C Si Mn S P Cr | CRM 15259 CRM 187A CRM 1761 CRM 187A CRM BS14B CRM 15259 | 0.008217 0.008255 0.005601 0.000996 0.001466 0.007566 |

4- Calculation of expanded uncertainty

Uncertainty budget or expanded uncertainty calculated according to ISO/IEC 17025-2005 multiplied by a coverage factor k=2, providing a level of confidence of approximately 95%. Expanded uncertainty equal combined uncertainty multiplied by 2. Take a root mean square of summation of all combined uncertainty for every budget of uncertainty after taking a suitable distribution with corresponding divisor, e.g calculation of repeatability (run to run variation) used normal distribution with divisor 1, calculation of CRM uncertainty used normal distribution with divisor 2, calculation of curve standard error and bias of standard deviation used rectangular distribution with divisor 1.732051 and in combined, expanded uncertainty calculation used normal distribution with divisor 1

Table (4) : Calculated Uncertainty for each element

| Element | Concentration | Uncertainty ±U |
| --- | --- | --- |
| C Si Mn S P Cr | 0.67 0.567 0.678 0.018 0.028 0.47 | 0.0334 0.0273 0.0178 0.0045 0.0032 0.0379 |

##### 5-Reporting

After uncertainty calculation the value of element percent write plus minus value of uncertainty followed by these statement (the reported expanded uncertainties are based on a standard uncertainties multiplied by a coverage factor k=2 providing a level of confidence of approximately 95%).

###### Proficiency test program

According to proficiency test program report. Test results for each element are grouped in the table according to the ASTM method utilized. Individual method summaries are provided at the end of each grouping enabling to compare laboratory data to the robust mean obtained by a given method. An overall summary robust mean combining all methods is provided at the end of the table for comparison to the grand mean. Means and standard deviations are computed using two stages. Robust procedure which- designed to limit the influence of the unusually large as small results while eliminating only the most egregious among them. The computations and philosophy of robust statistics are explained more fully in robust statistics – How not to reject outlier(5). Table (5) show robust mean, robust standard deviation and z score according to ASTM PT report

Table (5): Proficiency Test Result

| Element | Measured Element | Robust Mean | Robust StDev | z score |
| --- | --- | --- | --- | --- |
| C Si Mn S P Cr | 0.4999 0.55305 0.75576 0.01782 0.02269 0.505 | 0.4932 0.5601 0.7469 0.0191 0.0225 0.4938 | 0.0091 0.0118 0.0085 0.0014 0.0016 0.0082 | 0.8 -0.6 1 - 1.2 0.2 1.2 |

##### Conclusion

By comparison nominal uncertainty of CRM (from certificate) with estimated uncertainty in laboratory can obtained that, the percent of uncertainty for every element related to concentration of the same element is 4.7,4.05,0.7,5.5,6.4% for C, Si, Mn, S, P, Cr respectively, after calculated uncertainty in the condition of laboratory (estimated uncertainty) the previous percent changes to 4.9, 4.8, 2.6, 2.5, 11,8% for C, Si, Mn, S, P, Cr respectively, the difference between percent of nominal and estimated uncertainty is 0.43, 0.75, 1.9, 3, 3.9, 1.6 for the previous element respectively.

When compared z factor value obtained in proficiency test program of ASTM with consideration the sample used not similar in concentration but near as possible. Low difference between nominal and estimated uncertainty give low z score and vice versa

##### Acknowledgment

Thanks for all labor chemists, technician in ARCOSTEEL lab.

References: -

1- Taverniers I, Van Bockstaele E, De loose M, Trends Anal chem. 23; 480-490.

2-  Juniper IR (1999) Accred Qual Assur 4:336-341

3- Wood R, Thompson M (1996) Accred Qual Assur 1:49-56

4- Eurachem/Citac,Guide (4) Eurachem/Citac Guide, Quantifying uncertainty in analytical measurement 2nd ed, 2000

5- Analytical Methods Committee, 114, pp, 1693-97 Dec –1989

- ARL (applied research laboratory) software Documentation  WinOES V2.1 AA83395 multivariable regression (Switzerland) 1998.

Dear Prof. Dr. Paul De Biève

Independent Consultant on MIC

Former Advisor to the Director

Institute for Reference Materials and Measurement

I have pleasure in enclosing herewith article minscript in title “Uncertainty and Proficiency Testing Programs of Measurement of the Analysis Carbon, Silicon, Manganese and Chrome in Low Alloy Steel by Optical Emission Spectroscopy (OES), Calibrated with Certified Reference Materials” by Elmaghrabi A.M for publication in your journal “Journal for Quality, Comparability and Reliability in Chemical Measurement”.

I appreciate receiving a report on this paper stating whether it was accepted for publication or any modification need, please accept my regards and I look forward to receiving your reply as soon as possible.

Thank you for your kindness

Your sincerely

ElMaghrabi A.M

ARCOSTEEL Lab Mgr.

Sadat City – 4th industrial zone

P.O. 11 – Egypt

Email : gmaghraby@yahoo.com

Tel : +2010101983932